THEMIS EARTH FIELD COMPARISON

First Results

THM-SOC_112_ATTPROC

August 01, 2007

Dr. Uli Auster, THEMIS FGM

Dr. Krishan Khurana, THEMIS Science ACS

Michael Ludlam, THEMIS ISE

Vassilis Angelopoulos, THEMIS Principal Investigator

# Document Revision Record

| Rev. | Date | Description of Change | Approved By |
| --- | --- | --- | --- |
| 1 | 2007-08-01 | Draft | - |

# Distribution List

| Name | Email |
| --- | --- |
| Jim Lewis, U.C. Berkeley | jwl@ssl.berkeley.edu |
| P. R. Harvey, UCB, FSW | prh@ssl.berkeley.edu |
| Dr. Tai Phan, U.C. Berkeley | phan@ssl.berkeley.edu |
| Dr. Forrest Mozer, UCB | fmozer@ssl.berkeley.edu |
| Dr. Robert Ergun, LASP | ree@fast.colorado.edu |
| Dr. James McFadden. ESA | mcfadden@ssl.berkeley.edu |
| Dr. Chris Cully, LASP | cully@colorado.edu |
| Dr. Manfred Bester, MOM | Manfred@ssl.berkeley.edu |
| Dr. Ellen Taylor, U.C. Berkeley | ertaylor@ssl.berkeley.edu |
| Dr. Davin Larson, SST | davin@ssl.berkeley.edu |
| Dr. John Bonnell, EFI | jbonnell@ssl.berkeley.edu |
| Dr. Dave Sibeck, NASA GSFC | david.g.sibeck@nasa.gov |
| Dr. Harald Frey, UCB | hfrey@ssl.berkeley.edu |
| Daniel Cosgrove, UCB, MOC-ACS | dcosgrove@ssl.berkeley.edu |

# TBD List

| Identifier | Description |
| --- | --- |

Table of Contents

Document Revision Record	2

Distribution List	2

TBD List	2

1.	Introduction	4

1.1	Purpose and Scope.	4

1.2	Applicable Documents.	4

2.	First Results	4

2.1	Summary of Results	5

2.1.1	2007, Mar 04 to 06	5

2.1.2	2007, June 08	5

2.1.3        2007, June 29	5

# Introduction

## Purpose and Scope.

Verification of THEMIS attitude and magnetic field measurements by comparison with the Earth’s magnetic field.

## Applicable Documents.

- THM_SYS_012_PDMP				THEMIS Project Data Management Plan

- THM_SOC_101_TIME				THEMIS TIME Definition

- THM_SOC_108_GMAG_L1_VARNAMES		THEMIS GMAG Variable Name Def’s

- THM_SOC_111_SUNSENSPROC			THEMIS SUN SENSOR Science Processing

- THM_SOC_112_SCIATT_PROC			THEMIS Science ATT Determination

- THM_SOC_113_FGM_CALPROC		THEMIS FGM CAL File and Processing

- THM_SOC_114_SCM_CALPROC			THEMIS SCM CAL File and Processing

- THM_SOC_115_EFI_CALPROC			THEMIS EFI CAL File and Processing

- THM_SOC_116_ESA_CALPROC			THEMIS ESA CAL File and Processing

- THM_SOC_117_SST_CALPROC			THEMIS SST CAL File and Processing

- SAI-SPEC-1079A (Oct 26, 2005)			THEMIS Coordinate systems

- SAI-RPT-0722a (September, 2006)			Probe Alignment Report (MSSS data, p18)

- pturin e-mail on Faro alignment results (9/28/06)	FGM, SCM mag alignments

- THM-MB-MEC-005-Magnetometer clocking r7.pdf 	MAG clocking angles

# First Results

We performed our first Earth field comparisons at times when we had the luxury of having both fgl and fge data at perigee for all five spacecraft (2007, Mar 04 and 06). The next fits we did for 2007, June 08 when we again had both fgl and (improved) fge data from THEMIS C, D and E. We did another set of fits for 2007, June 29. Our fits contained data inside of 2 Re. We used the Tsyganenko 2001 as well as the IGRF 2005 models. In order to check how stable the results are we divided each pass into two parts and calculated the parameters for the first half, the second half and for the whole pass. We did not use times when the magnetometers were in saturation or when the spacecraft were eclipsed. When the spacecraft are eclipsed the phase angle shows very strong drifts. Figure 1 shows such an example. Our new software is able to fix such drifts in phi.

Figure 1. THEMIS C data and model are displayed in DSL-coordinates. The gray region marks the eclipsed period. The two spin plane components show large differences to the model due to drifting phi angle.

Calibration parameters that we solve for:

Gxy	: gain correction of the spin plane gains (multiplies previous gain)

Gz	: gain correction of the spin axis gain (multiplies previous gain)

dphi	: zero angle correction of the spin plane (subtract from existing phi)

dras	: correction to right ascension of the true spin axis (add to existing ras)

ddec	: correction to declination of the true spin axis (add to existing dec)

We found that the phase is drifting for all 5 spacecraft at perigee even if the spacecraft are not eclipsed. This caused problems for our software that we had developed before the launch. We had to rewrite the software so that all calibration parameters can be determined without knowing the precise phase angle. Furthermore we were able to calculate dphi in high time resolution. Solving for dphi in high time resolution greatly improved the RMS error of our fits. Future reports will also contain the other parameters in higher time resolution which appear to be more stable than dphi.

The values that we calculated for 2007, June 08 and June 29 (‘full flyby’) represent our currently best results and can be used to improve attitude and magnetometer calibration. Nevertheless we need to further investigate the accuracy of our calculations especially dphi. Tables 20 and 21 give a summary of the first results.

## Summary of Results

### 2007, Mar 04 to 06

The RMS errors of the fits that we did using fge data are higher than those we got from fgl data. Most calibration parameters that we got from fge data are roughly in agreement with the ones from fgl data except for dphi. dphi is significantly different between fge and fgl data. Most results that we got for Mar 04 are not stable. We get different results for “first half”, “full flyby” and “second half” (see Tables 1a, 1b, 5a, 5b, 9a, 9b, 13a, 13b). On Mar 06 the results are more stable (see Tables 2a, 2b, 6a, 6b, 10a, 10b, 14a, 14b, 17a, 17b).

### 2007, June 08

On June 08 THEMIS C, D and E again have simultaneous fge and fgl data at perigee. The fge data had been greatly improved since Mar 06. The fits for fge and fgl data give the same results and the RMS of fit is also roughly the same. This means that fge data can now be used for Earth field comparison. dphi is drifting for all 5 spacecraft (see Figures 2, 3, 4, 5 and 6).

Figure 2. Themis A, drifting phase after eclipse.

Figure 3. Themis B, drifting phase after eclipse.

Figure 4. Themis C, drifting phase after eclipse.

Figure 5. Themis D, drifting phase after eclipse.

Figure 6. Themis E, drifting phase after eclipse.

Having dphi in high resolution drastically improves the RMS of fit of the spin plane components x and y (see Tables 3a, 3b, 7a, 7b, 11a, 11b, 15a, 15b, 18a, 18b).

2.1.3 2007, June 29

The results for June 29 are similar to the results from June 08. dphi again is drifting for all 5 spacecraft (see Figures 7, 8, 9, 10 and 11).

Figure 7. Themis A, drifting phase after eclipse.

Figure 8. Themis B, drifting phase after eclipse.

Figure 9. Themis C, drifting phase after eclipse.

Figure 10. Themis D, drifting phase after eclipse.

Figure 11. Themis E, drifting phase after eclipse.

Table 1a

2007, Mar 04

THEMIS A fgl

RMS x in [nT]:              54.6         123.2         165.5

RMS y in [nT]:              89.3          83.7          77.7

RMS z in [nT]:              66.3          59.4          51.5

first half    full flyby    second half

Gxy          :         0.9965189     0.9966977     0.9971170

Gz           :         0.9944776     0.9941838     0.9947261

phi     [deg]:              1.43          0.71          1.02

ras     [deg]:             -0.62         -0.53         -0.42

dec     [deg]:             -0.37         -0.22         -0.19

RMS x out[nT]:              32.8          69.0          35.7

RMS y out[nT]:              71.3          49.4          27.6

RMS z out[nT]:               4.5           9.3           2.0

high res phi:

RMS x out[nT]; full flyby:                15.0

RMS y out[nT]; full flyby:                21.9

Table 1b

THEMIS A fge

RMS x in [nT]:              75.7          98.7         117.3

RMS y in [nT]:             123.9         107.6          88.4

RMS z in [nT]:              66.2          59.4          51.6

first half    full flyby    second half

Gxy          :         0.9963310     0.9967919     0.9972618

Gz           :         0.9944870     0.9941942     0.9946720

phi     [deg]:              0.73         -0.06          0.42

ras     [deg]:             -0.61         -0.53         -0.42

dec     [deg]:             -0.37         -0.22         -0.19

RMS x out[nT]:              48.8         120.3         106.1

RMS y out[nT]:              86.5          81.9          78.8

RMS z out[nT]:               4.5           9.3           2.5

high res phi:

RMS x out[nT]; full flyby:                26.0

RMS y out[nT]; full flyby:                34.8

Table 2a

2007, Mar 06

THEMIS A fgl

RMS x in [nT]:              30.8          78.6         106.7

RMS y in [nT]:              85.4          66.9          40.8

RMS z in [nT]:              41.1          51.9          60.7

first half    full flyby    second half

Gxy          :         0.9989524     0.9998360     0.9999474

Gz           :         0.9977012     0.9980796     0.9997207

phi     [deg]:              1.00          0.60          0.65

ras     [deg]:             -0.20         -0.12         -0.02

dec     [deg]:             -0.26         -0.25         -0.29

RMS x out[nT]:              31.9          19.8          22.8

RMS y out[nT]:              55.7          23.0          16.2

RMS z out[nT]:               5.3           9.4           2.8

high res phi:

RMS x out[nT]; full flyby:                 8.8

RMS y out[nT]; full flyby:                13.0

Table 2b

THEMIS A fge

RMS x in [nT]:              58.2          72.2          83.9

RMS y in [nT]:             131.6         100.6          54.1

RMS z in [nT]:              40.7          51.6          60.6

first half    full flyby    second half

Gxy          :         0.9989430     0.9997783     0.9998667

Gz           :         0.9977323     0.9981029     0.9996478

phi     [deg]:              0.54         -0.17         -0.14

ras     [deg]:             -0.19         -0.12         -0.02

dec     [deg]:             -0.26         -0.25         -0.29

RMS x out[nT]:              68.6          53.7          62.8

RMS y out[nT]:             107.0          49.8          44.4

RMS z out[nT]:               5.4           9.2           2.3

high res phi:

RMS x out[nT]; full flyby:                21.3

RMS y out[nT]; full flyby:                29.8

Table 3

2007, June 08

THEMIS A fge

RMS x in [nT]:              44.8          41.7          38.2

RMS y in [nT]:              97.8          70.4          18.8

RMS z in [nT]:               9.0           6.9           3.8

first half    full flyby    second half

Gxy          :         0.9981498     0.9993067     0.9984987

Gz           :         0.9987093     0.9990422     0.9968192

phi     [deg]:              0.85          0.79          0.51

ras     [deg]:              0.00          0.09          0.20

dec     [deg]:             -0.10         -0.08         -0.17

RMS x out[nT]:              14.2          20.5           9.2

RMS y out[nT]:              25.5          18.2           6.0

RMS z out[nT]:               1.2           1.1           0.7

high res phi:

RMS x out[nT]; full flyby:                 3.1

RMS y out[nT]; full flyby:                 3.5

Comment: ‘second half’ has lower field values so parameters are not as well determined as ‘first half’ and ‘full flyby’

Table 4

2007, June 29

THEMIS A fge

RMS x in [nT]:              46.4          38.2          27.5

RMS y in [nT]:             201.3         146.3          47.5

RMS z in [nT]:               6.8           5.5           3.8

first half    full flyby    second half

Gxy          :         0.9972020     0.9973828     0.9965912

Gz           :         1.0000983     1.0005693     1.0005354

phi     [deg]:              1.20          1.09          0.55

ras     [deg]:             -0.22         -0.20         -0.19

dec     [deg]:             -0.05         -0.02         -0.02

RMS x out[nT]:              22.9          39.3          16.0

RMS y out[nT]:              30.7          27.4           7.2

RMS z out[nT]:               0.7           0.7           0.5

high res phi:

RMS x out[nT]; full flyby:                 3.6

RMS y out[nT]; full flyby:                 3.3

Comment: ‘second half’ has lower field values so parameters are not as well determined as ‘first half’ and ‘full flyby’

Table 5a

2007, Mar 04

THEMIS B fgl

RMS x in [nT]:              58.2         144.3         195.5

RMS y in [nT]:              41.4          89.7         119.9

RMS z in [nT]:             141.0         162.8         182.1

first half    full flyby    second half

Gxy          :         1.0043737     1.0031528     1.0008520

Gz           :         0.9925750     1.0021864     1.0022320

phi     [deg]:             -0.28         -0.05         -0.61

ras     [deg]:             -2.39         -2.79         -3.89

dec     [deg]:              0.31          0.37          0.13

RMS x out[nT]:              65.9          26.8          38.5

RMS y out[nT]:              12.2          29.8          28.8

RMS z out[nT]:               5.0          28.8           8.8

high res phi:

RMS x out[nT]; full flyby:                 9.4

RMS y out[nT]; full flyby:                23.6

Table 5b

THEMIS B fge

RMS x in [nT]:             114.3         140.9         163.2

RMS y in [nT]:              95.9         122.5         144.3

RMS z in [nT]:             141.0         162.8         182.1

first half    full flyby    second half

Gxy          :         1.0042965     1.0031619     1.0010045

Gz           :         0.9925851     1.0021893     1.0022622

phi     [deg]:             -0.85         -0.37         -0.98

ras     [deg]:             -2.39         -2.79         -3.88

dec     [deg]:              0.31          0.37          0.13

RMS x out[nT]:             164.7         109.1         110.9

RMS y out[nT]:              68.0          71.4          61.5

RMS z out[nT]:               5.0          28.8           8.8

high res phi:

RMS x out[nT]; full flyby:                19.3

RMS y out[nT]; full flyby:                52.1

Table 6a

2007, Mar 06

THEMIS B fgl

RMS x in [nT]:             106.8          77.6          25.2

RMS y in [nT]:              45.0          55.7          64.6

RMS z in [nT]:              57.0          54.2          51.3

first half    full flyby    second half

Gxy          :         1.0001987     0.9998711     1.0018684

Gz           :         0.9982339     0.9978267     0.9987373

phi     [deg]:              0.40          0.48          0.59

ras     [deg]:              0.49          0.59          0.59

dec     [deg]:              0.22          0.22          0.35

RMS x out[nT]:              38.3          22.4          13.7

RMS y out[nT]:              18.0          18.3          17.1

RMS z out[nT]:               5.1           5.3           1.8

high res phi:

RMS x out[nT]; full flyby:                 6.6

RMS y out[nT]; full flyby:                13.1

Table 6b

THEMIS B fge

RMS x in [nT]:             140.3         109.4          65.5

RMS y in [nT]:              94.6          84.8          73.7

RMS z in [nT]:              56.5          54.2          51.7

first half    full flyby    second half

Gxy          :         0.9996441     0.9993771     1.0020151

Gz           :         0.9981898     0.9977838     0.9987806

phi     [deg]:              0.19          0.10          0.19

ras     [deg]:              0.50          0.59          0.59

dec     [deg]:              0.22          0.22          0.37

RMS x out[nT]:             124.3          90.5          38.0

RMS y out[nT]:              87.7          70.9          48.2

RMS z out[nT]:               3.9           4.9           2.1

high res phi:

RMS x out[nT]; full flyby:                25.8

RMS y out[nT]; full flyby:                42.9

Table 7

2007, June 08

THEMIS B fge

RMS x in [nT]:             128.5         106.4          78.3

RMS y in [nT]:             204.5         146.3          31.3

RMS z in [nT]:              19.4          22.8          25.7

first half    full flyby    second half

Gxy          :         1.0005740     1.0006453     0.9995988

Gz           :         1.0000900     1.0000581     0.9966351

phi     [deg]:              1.03          0.96          0.71

ras     [deg]:             -0.64         -0.63         -0.42

dec     [deg]:              0.01          0.01         -0.06

RMS x out[nT]:              20.0          20.6           1.8

RMS y out[nT]:              15.1          12.4           2.5

RMS z out[nT]:               1.1           0.9           0.5

high res phi:

RMS x out[nT]; full flyby:                 1.2

RMS y out[nT]; full flyby:                 2.4

Comment: ‘second half’ has lower field values so parameters are not as well determined as ‘first half’ and ‘full flyby’

Table 8

2007, June 29

THEMIS B fge

RMS x in [nT]:             106.1          90.1          70.6

RMS y in [nT]:             207.9         150.3          44.7

RMS z in [nT]:              14.9          15.9          16.7

first half    full flyby    second half

Gxy          :         1.0000881     1.0000386     0.9994189

Gz           :         1.0008000     1.0008514     1.0000553

phi     [deg]:              0.94          0.90          0.72

ras     [deg]:             -0.48         -0.48         -0.41

dec     [deg]:             -0.07         -0.06         -0.05

RMS x out[nT]:              12.2          18.6           9.0

RMS y out[nT]:              10.1           8.4           2.5

RMS z out[nT]:               1.6           1.3           0.8

high res phi:

RMS x out[nT]; full flyby:                 1.9

RMS y out[nT]; full flyby:                 2.0

Comment: ‘second half’ has lower field values so parameters are not as well determined as ‘first half’ and ‘full flyby’

Table 9a

2007, Mar 04

THEMIS C fgl

RMS x in [nT]:              59.8         121.3         160.8

RMS y in [nT]:              53.8          49.2          44.1

RMS z in [nT]:              47.3          88.5         115.9

first half    full flyby    second half

Gxy          :         1.0009185     1.0010380     1.0012914

Gz           :         0.9993683     0.9992903     0.9998342

phi     [deg]:              0.70          0.74          0.89

ras     [deg]:              0.17          0.13          0.10

dec     [deg]:             -0.41         -0.43         -0.46

RMS x out[nT]:               8.1          30.0          30.5

RMS y out[nT]:              11.7          14.8          17.0

RMS z out[nT]:               2.7           4.4           3.3

high res phi:

RMS x out[nT]; full flyby:                 5.4

RMS y out[nT]; full flyby:                 3.8

Table 9b

THEMIS C fge

RMS x in [nT]:              48.9         101.8         135.5

RMS y in [nT]:              97.4          95.7          94.0

RMS z in [nT]:              47.3          88.5         115.9

first half    full flyby    second half

Gxy          :         1.0009838     1.0011536     1.0014312

Gz           :         0.9993728     0.9992970     0.9998780

phi     [deg]:              0.18          0.12          0.46

ras     [deg]:              0.17          0.13          0.11

dec     [deg]:             -0.41         -0.43         -0.46

RMS x out[nT]:              37.0         101.9         121.6

RMS y out[nT]:              41.4          68.6          86.3

RMS z out[nT]:               2.5           4.4           3.3

high res phi:

RMS x out[nT]; full flyby:                30.6

RMS y out[nT]; full flyby:                38.1

Table 10a

2007, Mar 06

THEMIS C fgl

RMS x in [nT]:              69.3         105.8         132.6

RMS y in [nT]:             146.8         149.6         152.4

RMS z in [nT]:             116.8          89.7          49.3

first half    full flyby    second half

Gxy          :         1.0017283     1.0015061     1.0009347

Gz           :         0.9992228     0.9990990     0.9987689

phi     [deg]:              0.44          0.56          0.64

ras     [deg]:             -1.34         -1.33         -1.34

dec     [deg]:             -0.46         -0.45         -0.47

RMS x out[nT]:              28.4          16.5          13.2

RMS y out[nT]:              13.5          17.1          15.3

RMS z out[nT]:               4.0           4.2           4.2

high res phi:

RMS x out[nT]; full flyby:                 4.2

RMS y out[nT]; full flyby:                 4.6

Table 10b

THEMIS C fge

RMS x in [nT]:             111.8         110.9         110.0

RMS y in [nT]:             154.2         145.9         137.2

RMS z in [nT]:             117.1          89.9          49.4

first half    full flyby    second half

Gxy          :         1.0017282     1.0015631     1.0009624

Gz           :         0.9991940     0.9990796     0.9987004

phi     [deg]:             -0.03          0.14          0.26

ras     [deg]:             -1.34         -1.33         -1.34

dec     [deg]:             -0.46         -0.45         -0.48

RMS x out[nT]:              85.5          71.7          59.2

RMS y out[nT]:              63.1          56.3          49.3

RMS z out[nT]:               5.1           5.0           4.7

high res phi:

RMS x out[nT]; full flyby:                22.8

RMS y out[nT]; full flyby:                30.9

Table 11a

2007, June 08

THEMIS C fgl

RMS x in [nT]:              68.3          61.6          54.0

RMS y in [nT]:              69.5          49.8          11.5

RMS z in [nT]:              18.4          14.0           7.1

first half    full flyby    second half

Gxy          :         1.0005465     1.0004874     1.0025643

Gz           :         0.9999694     0.9998866     1.0033844

phi     [deg]:              0.51          0.51          0.48

ras     [deg]:              0.03          0.02         -0.19

dec     [deg]:             -0.11         -0.11          0.00

RMS x out[nT]:              11.2          11.1          11.3

RMS y out[nT]:              19.0          13.8           4.7

RMS z out[nT]:               1.4           1.0           0.5

high res phi:

RMS x out[nT]; full flyby:                 1.5

RMS y out[nT]; full flyby:                 2.2

Comment: ‘second half’ has lower field values so parameters are not as well determined as ‘first half’ and ‘full flyby’

Table 11b

THEMIS C fge

RMS x in [nT]:              68.4          61.6          54.0

RMS y in [nT]:              69.4          49.7          11.5

RMS z in [nT]:              18.4          14.0           7.1

first half    full flyby    second half

Gxy          :         1.0006871     1.0005327     1.0029202

Gz           :         1.0000637     0.9998952     1.0040491

phi     [deg]:              0.51          0.51          0.47

ras     [deg]:              0.03          0.02         -0.23

dec     [deg]:             -0.10         -0.11          0.01

RMS x out[nT]:              11.2          11.1          11.3

RMS y out[nT]:              19.1          13.9           4.7

RMS z out[nT]:               1.4           1.1           0.5

high res phi:

RMS x out[nT]; full flyby:                 1.5

RMS y out[nT]; full flyby:                 2.2

Comment: ‘second half’ has lower field values so parameters are not as well determined as ‘first half’ and ‘full flyby’

Table 12

2007, June 29

THEMIS C fge

RMS x in [nT]:              70.8          62.5          52.9

RMS y in [nT]:              41.7          30.5          11.1

RMS z in [nT]:              60.7          47.5          28.8

first half    full flyby    second half

Gxy          :         0.9997712     0.9994029     0.9997386

Gz           :         1.0007656     1.0002378     1.0014532

phi     [deg]:              0.40          0.41          0.39

ras     [deg]:              0.35          0.35          0.24

dec     [deg]:             -0.25         -0.28         -0.27

RMS x out[nT]:              12.6          11.6          11.3

RMS y out[nT]:              19.5          14.0           5.5

RMS z out[nT]:               1.0           0.8           0.6

high res phi:

RMS x out[nT]; full flyby:                 2.3

RMS y out[nT]; full flyby:                 2.4

Comment: ‘second half’ has lower field values so parameters are not as well determined as ‘first half’ and ‘full flyby’

Table 13a

2007, Mar 04

THEMIS D fgl

RMS x in [nT]:              70.1         156.4         209.8

RMS y in [nT]:              82.5          81.4          80.4

RMS z in [nT]:              37.8          49.0          58.1

first half    full flyby    second half

Gxy          :         0.9994797     0.9991542     0.9985943

Gz           :         0.9980148     0.9976051     0.9950979

phi     [deg]:              1.06          1.04          1.06

ras     [deg]:             -0.12         -0.15         -0.24

dec     [deg]:             -0.29         -0.24         -0.19

RMS x out[nT]:              43.3          45.2          45.5

RMS y out[nT]:              41.8          36.9          31.9

RMS z out[nT]:               3.9           5.5           4.5

high res phi:

RMS x out[nT]; full flyby:                25.7

RMS y out[nT]; full flyby:                26.5

Table 13b

THEMIS D fge

RMS x in [nT]:              63.8         103.6         131.9

RMS y in [nT]:              84.9          87.3          89.7

RMS z in [nT]:              37.8          49.0          58.1

first half    full flyby    second half

Gxy          :         0.9996476     0.9991661     0.9985691

Gz           :         0.9980176     0.9976066     0.9950850

phi     [deg]:              0.50          0.65          0.55

ras     [deg]:             -0.12         -0.15         -0.24

dec     [deg]:             -0.29         -0.24         -0.19

RMS x out[nT]:              62.2          83.9          98.8

RMS y out[nT]:              70.9          76.3          83.1

RMS z out[nT]:               3.8           5.4           4.4

high res phi:

RMS x out[nT]; full flyby:                35.8

RMS y out[nT]; full flyby:                47.3

Table 14a

2007, Mar 06

THEMIS D fgl

RMS x in [nT]:              95.2         119.6         139.8

RMS y in [nT]:             106.4          85.1          56.2

RMS z in [nT]:              49.2          50.1          51.0

first half    full flyby    second half

Gxy          :         0.9988832     0.9994855     0.9997411

Gz           :         0.9982475     0.9983416     0.9985512

phi     [deg]:              1.13          0.82          0.73

ras     [deg]:             -0.05          0.02          0.13

dec     [deg]:             -0.26         -0.22         -0.23

RMS x out[nT]:              29.0          16.3           9.8

RMS y out[nT]:              41.1          26.4           6.8

RMS z out[nT]:               5.1           7.8           2.1

high res phi:

RMS x out[nT]; full flyby:                 7.6

RMS y out[nT]; full flyby:                12.7

Table 14b

THEMIS D fge

RMS x in [nT]:              62.0          89.5         110.4

RMS y in [nT]:             103.9          86.2          63.7

RMS z in [nT]:              48.9          49.9          50.8

first half    full flyby    second half

Gxy          :         0.9989533     0.9994563     0.9996572

Gz           :         0.9982728     0.9983698     0.9986472

phi     [deg]:              0.48          0.26          0.25

ras     [deg]:             -0.05          0.02          0.13

dec     [deg]:             -0.26         -0.22         -0.23

RMS x out[nT]:              60.5          82.7         102.5

RMS y out[nT]:              69.6          66.6          60.8

RMS z out[nT]:               5.3           7.9           2.8

high res phi:

RMS x out[nT]; full flyby:                28.0

RMS y out[nT]; full flyby:                44.4

Table 15a

2007, June 08

THEMIS D fgl

RMS x in [nT]:              63.3          61.0          58.5

RMS y in [nT]:              87.5          63.2          18.8

RMS z in [nT]:              13.0           9.4           2.7

first half    full flyby    second half

Gxy          :         0.9992955     0.9994763     0.9997749

Gz           :         0.9996458     0.9996917     0.9994259

phi     [deg]:              0.53          0.54          0.51

ras     [deg]:             -0.09         -0.09         -0.07

dec     [deg]:             -0.10         -0.10         -0.10

RMS x out[nT]:              12.6          13.1          14.0

RMS y out[nT]:              19.9          14.6           5.1

RMS z out[nT]:               1.7           1.5           1.3

high res phi:

RMS x out[nT]; full flyby:                 2.1

RMS y out[nT]; full flyby:                 3.0

Comment: ‘second half’ has lower field values so parameters are not as well determined as ‘first half’ and ‘full flyby’

Table 15b

THEMIS D fge

RMS x in [nT]:              63.4          61.0          58.5

RMS y in [nT]:              87.4          63.2          18.7

RMS z in [nT]:              13.0           9.4           2.7

first half    full flyby    second half

Gxy          :         0.9994534     0.9994921     1.0004619

Gz           :         0.9997641     0.9996944     1.0008086

phi     [deg]:              0.54          0.54          0.51

ras     [deg]:             -0.08         -0.09         -0.14

dec     [deg]:             -0.09         -0.10         -0.06

RMS x out[nT]:              12.6          13.2          14.3

RMS y out[nT]:              19.9          14.7           5.2

RMS z out[nT]:               1.7           1.5           1.3

high res phi:

RMS x out[nT]; full flyby:                 2.1

RMS y out[nT]; full flyby:                 3.0

Comment: ‘second half’ has lower field values so parameters are not as well determined as ‘first half’ and ‘full flyby’

Table 16

2007, June 29

THEMIS D fge

RMS x in [nT]:              70.5          62.8          53.9

RMS y in [nT]:              64.7          47.8          19.4

RMS z in [nT]:              52.7          40.2          21.5

first half    full flyby    second half

Gxy          :         0.9991764     0.9984543     0.9997635

Gz           :         1.0007696     0.9994311     1.0025448

phi     [deg]:              0.45          0.45          0.37

ras     [deg]:              0.17          0.17         -0.11

dec     [deg]:             -0.21         -0.28         -0.30

RMS x out[nT]:              21.5          17.8          13.8

RMS y out[nT]:              25.5          19.0           9.7

RMS z out[nT]:              23.0          18.5          12.3

high res phi:

RMS x out[nT]; full flyby:                10.6

RMS y out[nT]; full flyby:                 8.3

Comment: ‘second half’ has lower field values so parameters are not as well determined as ‘first half’ and ‘full flyby’

Table 17a

2007, Mar 06

THEMIS E fgl

RMS x in [nT]:              70.8         156.0         208.9

RMS y in [nT]:             107.3          91.0          71.0

RMS z in [nT]:              43.9          56.5          66.7

first half    full flyby    second half

Gxy          :         0.9967442     0.9988005     0.9990300

Gz           :         0.9983096     0.9990097     1.0008110

phi     [deg]:              1.16          0.92          1.13

ras     [deg]:             -0.13         -0.12         -0.09

dec     [deg]:             -0.22         -0.31         -0.32

RMS x out[nT]:              49.1          45.0          45.3

RMS y out[nT]:              39.0          24.0          28.2

RMS z out[nT]:               5.3           9.4           9.9

high res phi:

RMS x out[nT]; full flyby:                 6.7

RMS y out[nT]; full flyby:                21.9

Table 17b

THEMIS E fge

RMS x in [nT]:              75.2         131.2         169.6

RMS y in [nT]:             119.1         100.8          78.4

RMS z in [nT]:              43.5          54.1          62.8

first half    full flyby    second half

Gxy          :         0.9966804     0.9991137     0.9997218

Gz           :         0.9983225     0.9989001     1.0006769

phi     [deg]:              0.66          0.52          0.73

ras     [deg]:             -0.12         -0.13         -0.12

dec     [deg]:             -0.21         -0.29         -0.30

RMS x out[nT]:              80.8         102.1         118.5

RMS y out[nT]:              73.8          69.3          66.1

RMS z out[nT]:               5.4           6.0           3.3

high res phi:

RMS x out[nT]; full flyby:                24.3

RMS y out[nT]; full flyby:                48.3

Table 18a

2007, June 08

THEMIS E fgl

RMS x in [nT]:              70.3          65.9          61.2

RMS y in [nT]:             106.7          77.0          22.0

RMS z in [nT]:              16.6          12.4           5.6

first half    full flyby    second half

Gxy          :         0.9994212     0.9996015     0.9993776

Gz           :         1.0003965     1.0005305     1.0003970

phi     [deg]:              0.66          0.67          0.65

ras     [deg]:             -0.03          0.00          0.03

dec     [deg]:             -0.07         -0.06         -0.05

RMS x out[nT]:              12.0          11.9          11.2

RMS y out[nT]:              20.0          14.6           4.9

RMS z out[nT]:               1.7           1.4           1.1

high res phi:

RMS x out[nT]; full flyby:                 2.0

RMS y out[nT]; full flyby:                 2.4

Comment: ‘second half’ has lower field values so parameters are not as well determined as ‘first half’ and ‘full flyby’

Table 18b

THEMIS E fge

RMS x in [nT]:              70.3          65.9          61.0

RMS y in [nT]:             106.5          76.9          21.8

RMS z in [nT]:              16.6          12.4           5.6

first half    full flyby    second half

Gxy          :         0.9994446     0.9996150     0.9979574

Gz           :         1.0003477     1.0005733     0.9978671

phi     [deg]:              0.66          0.67          0.66

ras     [deg]:             -0.03          0.00          0.19

dec     [deg]:             -0.07         -0.06         -0.12

RMS x out[nT]:              11.9          11.8          11.2

RMS y out[nT]:              20.0          14.6           4.8

RMS z out[nT]:               1.7           1.4           1.1

high res phi:

RMS x out[nT]; full flyby:                 2.0

RMS y out[nT]; full flyby:                 2.4

Comment: ‘second half’ has lower field values so parameters are not as well determined as ‘first half’ and ‘full flyby’

Table 19

2007, June 29

THEMIS E fge

RMS x in [nT]:              79.5          69.5          57.8

RMS y in [nT]:              85.5          62.8          24.3

RMS z in [nT]:              48.7          37.3          20.3

first half    full flyby    second half

Gxy          :         0.9984340     0.9992235     0.9963685

Gz           :         0.9997738     1.0015388     0.9970961

phi     [deg]:              0.55          0.59          0.63

ras     [deg]:              0.14          0.16          0.50

dec     [deg]:             -0.27         -0.18         -0.23

RMS x out[nT]:              12.2          11.5          10.4

RMS y out[nT]:              18.3          14.1           4.3

RMS z out[nT]:               1.7           1.4           0.4

high res phi:

RMS x out[nT]; full flyby:                 3.9

RMS y out[nT]; full flyby:                 3.9

Comment: ‘second half’ has lower field values so parameters are not as well determined as ‘first half’ and ‘full flyby’

Table 20

| 2007, Jun 08 | Gxy | Gz | dphi [deg] | dras [deg] | ddec [deg] |
| --- | --- | --- | --- | --- | --- |
| THEMIS A | 0.9993067 | 0.9990422 | 0.79 | 0.09 | -0.08 |
| THEMIS B | 1.0006453 | 1.0000581 | 0.96 | -0.63 | 0.01 |
| THEMIS C | 1.0004874 | 0.9998866 | 0.51 | 0.02 | -0.11 |
| THEMIS D | 0.9994763 | 0.9996917 | 0.54 | -0.09 | -0.10 |
| THEMIS E | 0.9996015 | 1.0005305 | 0.67 | 0.00 | -0.06 |

Please note:  dphi should be subtracted from existing phase

Table 21

| 2007, Jun 29 | Gxy | Gz | dphi [deg] | dras [deg] | ddec [deg] |
| --- | --- | --- | --- | --- | --- |
| THEMIS A | 0.9973828 | 1.0005693 | 1.09 | -0.20 | -0.02 |
| THEMIS B | 1.0000386 | 1.0008514 | 0.90 | -0.48 | -0.06 |
| THEMIS C | 0.9994029 | 1.0002378 | 0.41 | 0.35 | -0.28 |
| THEMIS D | 0.9984543 | 0.9994311 | 0.45 | 0.17 | -0.28 |
| THEMIS E | 0.9992235 | 1.0015388 | 0.59 | 0.16 | -0.18 |

Please note:  dphi should be subtracted from existing phase